University of Cambridge

Department of Materials Science and Metallurgy

X-ray Laboratory

Suggestions for Texture or Orientation Work

1. Transmission - Can you get the beam though it? If yes,

1.1 Take a flat plate photo in transmission.

Record sample position + mark top of film.

May need to take photo down 2 of the 3 ref directions.

(Could be quantified by collecting on Nanostar or with CCD and doing azimuthal scan calc).

Can also collect quantitative data in transmission on B2 – scan phi to give azimuthal scan

Usually OK for polymers - calculate absorption coefficient, cut sample thinner, use harder radiation

1.2 For some samples may help to collect diffractometer data in reflection and transmission and compare the relative intensities.

2. Reflection - If you cannot get the beam though the sample you must work in reflection.

2.1 Do a normal 2θ/θ (2θ/ω) scan and compare intensities with reference data.

(May fail to get any data at all if the sample is highly oriented).

Look for systematic differences between families of reflections, eg (00l)s strong and (hk0)s weak implies that the vectors s00l are oriented normal to the plane of the sample, ie (00l) planes parallel to plane of sample. This could tell you that the sample is oriented and which vector is normal to the sample plane. Weak to medium levels of orientation can be quantified with the Rietveld orientation parameter. (Beware that intensity differences could come from experimental errors or changes in atomic coordinates.)

2.2 To measure out-of-plane orientation, do an omega () scan at fixed 2 with a smallish (1/12 or less) divergence slit. (If you have say 0.5 div slit the instrumental broadening is ~0.5 in , thus reduce slit to measure sharper (≤2) peaks. 0.5 slit is OK if you expect a peak width of say 8-10) This is also called a 'rocking curve' as the sample is rocked in omega.  The breadth (FWHM) of this peak is commonly used as a measure of out-of-plane orientation – works best for medium to strong orientation.

2.3 To measure in-plane orientation do a phi scan or measure a pole figure on an asymmetric peak (ie has an in-plane component - not normal to sample plane so has on ‘offset’ to the normal). The spread of these peaks give a measure of orientation. For some (highly crystalline) samples it might be possible to do in-plane measurements. Also om scans on skew symmetric peaks may give higher resolution than phi scans (because of line focus & way in which diffractometer constructed to be v accurate in 2t & om & less in psi and phi).

* There are spread sheets to calculate the offset (in psi or omega) for cubic, orthomobic and hexagonal systems, based on equations in Cullity (in Structures folder on all 3 standalone PCs)

(Several pole figures measured on different reflections may be required to describe a complex orientation. These can be converted from pole figures in reciprocal space to inverse pole figures in real space. This sort of work is not currently carried out in Mat Sci)

2.4 Can do various different maps on B2 or HRXRD eg 2t/om against psi, 2t/om v om etc to try & find reflections & work out orientation

2.5 Laue back reflection photo with Mo radiation (no filter, ie ‘white’ radiation)

Good for highly oriented (single crystal) samples (see Cullity)

Use WSLAUE (Marin and Dieguez’s program)  to simulate data and compare with experiment

Alternatively try Laue Altas in library to compare experimental and simulated data.

2.6 Back-reflection, with CuK, might help if only slightly oriented.

(Method 2.6 only works if reflections can be measured at 2>90)

2.7 Ward-Wallace texture camera, CuK - works for unoriented + well oriented samples, gives orientation wrt sample normal, ie offset but no in-plane info

2.8 Rotation photo, CuK, in say Stubbins camera (if you know 1 axis) (Unlikely to use this approach now)

2.9 Pole figures of at least 2 different reflections (see 3.)

(can be a Catch 22, if the sample is highly oriented, need to know the orientation to be able to set phi and psi to do 2θ/θ scan to find peak max on which to collect pole figure, see 2.4 and hints for further info)

Philips texture only works in reflection and at higher psi (or alpha) there is a loss of intensity due to defocussing. The unit is on a spot focus so resolution in chi/pisi and om are approx the same – better than B2 in chi but much worse than B2 in om and 2t. Collect data with Data Collector software & draw pole figures with Panalytical Texture software (can also view simply in data viewer)

On B2 can collect data in both reflection and transmission. However this is set on a line source so resolution in ω, θ 2θ is good (~0.03) but in χ, chi, psi it is poor (~1). See instructions in Using Bruker diffractometers.

.3. Quantification of Orientation

Usually involves measuring breadth (FWHM - full width half max) of a given peak egs:

3.1 Reflection mode, out-of-plane orientation, FWHM of 'rocking curve' ie 2theta fixed, omega scanned – best for medium to very highly oriented samples & a simple type of orientation

3.2 Reflection mode, in-plane orientation, FWHM of phi in scan or pole figure (psi & phi scanned, max at psi0) – most general method. Several well chosen pole figures are required to describe complex orientations. Highly oriented measure om in-plane or skew symmetric

3.3 Transmission mode, FWHM of a given azimuthal reflection

3.4 Rietveld refinement of data gives a single value for ‘orientation’ – works best for lower levels of a simple out-of-plane orientation.

# 4. Hints on setting up for Pole Figures

What is the phase that you wish to analyse?

Where do you expect peaks in 2theta?

Have you got a standard theta/2theta scan?

Is the sample oriented? Use one or more of the techniques in ‘Suggestions for Texture or Orientation Work’ to confirm that it is oriented and give you some clues about the type of orientation. Suggest 2.1, 2.4, 2.3 or 2.2 probably in this order.

You cannot rely on data from ICDD or another diffractometer to give the 2theta position. The sample height is so critical that you must determine the 2theta position each time the sample is put in the instrument.

Reflections that show up in a std theta/2theta scan can be measured in the texture goniometer at psi=0 and phi=0.  Pole figures can then be measured on these reflections.

If the sample is well oriented other reflections may be missing from the standard theta/2theta scan.

There are several different options to decide on where to set psi and phi in order to be able to do a theta/2theta scan (to determine 2theta at which to collect pole figure).

- Educated guess from what you expect, have found on other samples etc

- Trial and error

  - Find reflection with max at psi=0, eg pole figure on reflections observed in standard theta/2theta or 4 or 5 below

  - then calculate psi positions for other reflections (see below)

3.3 then determine phi and 2theta using optimise scan in Philips data collection software or manually

- Determine  from Ward Wallace photo and this gives psi then do 3.3

5    Use Laue photo to determine orientation of the flat face (can get this from std theta/2theta scan) and axes within it. Then calculate positions of other peaks (in phi and psi), knowing unit cell + orientation of sample wrt external axes.

Calc

This can be done in

1) spread sheets 2.3 or

2) is performed in Panalytical software (you set cell parameters & directions and ask for info on hkl - software calcs diff angles)]

Eg if the cell is cubic a=4.5 and (200) reflection max at psi=0 (and phi=0) angle  between the d1, (h00) planes,  and d2, (111) planes is

cosφ=h1h2+k1k2+l1l2(h12+k12+l12)(h22+k22+l22)

(All denominator is square root – see text books for equations for non cubic systems, cubic systems can also be calculated by simple geometry)

giving  =54.76 deg

Thus if (200) is at psi=0 (111) will be at psi=54.8

Still need to find phi, do phi scan - see 2.3.

Mary Vickers 1999

Modified Sept 05

Updated March 09

Again July09